Removing / inserting the Reading head

[Operation conditions]

Removing the reading head:

Perform this operation when it is necessary to remove the reading head.

|  | [Operation contents] - Remove the Slider holder. Remove both the contersunk head screws that secure the End cap. - Pull out the scale’s End cap. Note: Mounting insert and spring will jump out. |
| --- | --- |

- With tweezers lighlty peel off the adhesive on the dust lip’s.

- Partly remove the reading head and insert Mounting plate. With Securing plate you replace the measuring glass and fix the scanning unit in proper position.

Note:

If you do not insert Securing plate, some springs may jump out of the reading head!

Inserting the reading head:

- Adjust the scanning unit’s angle and lightly insert it. At this time, the scanning unit is set up to clasp the glass.

[Verification]  The position of the scanning unit must be parallel with the glass.

2) Following this, use the Slider holder to secure the reading head from moving.

Installing the End cap

[Operation conditions]

Perform this operation when putting together the scale unit after replacing the Reading head.

[Operation contents]

- Install the End cap seal (if removed) on the End cap, and lubricate it with the silicone grease.

- Insert the Mounting inserts and Spring 4 into the End cap.

[Note]	Be careful, that dents on the Mounting inserts must be oriented to the outside of the End cap.

- Push the End cap assembly into the scale.

- Secure the End cap assembly with 2 screws Ka25x6.

[Note]	Take the utmost care to not allow the entrance of dust and the like when wiping away filth from the scale side as it is very easy for it to enter at this time.