Service parts

TGM133 series disassembly drawing

Parts list

U  = a part is unavailable

| Scale’s main unit parts |
| --- |
| No. | Part name | Part code | Remarks | Ordering unit |
| 1 | Reading head | 430 263 020 001 | 0.1μm DS | 1 |
|  |  | 430 263 020 002 | 0.5μm DS | 1 |
|  |  | 430 263 020 003 | 1μm DS | 1 |
|  |  | 430 263 020 010 | 20μm/1Vp-p | 1 |
|  |  | 430 263 020 020 | 20μm/11Ap-p | 1 |
| 1a | Slider |  |  | U |
| 1b | Scanning unit |  |  | U |
| 2 | Cable | 430 263 030 XXX | SI or SV signals | 1 |
|  |  |  | Connector type |  |
|  |  |  | Cable length |  |
|  |  |  | Armor (with or without) |  |
|  |  | 430 263 031 XXX | DS signals | 1 |
|  |  |  | Connector type |  |
|  |  |  | Cable length |  |
|  |  |  | Armor (with or without) |  |
| 3 | Al housing assembly | 430 261 520 XXX | XXX = Lm (cm) | U |
| 4 | End cap – right | 430 261 180 100 | Assembly | 10 |
| 5 | End cap – left | 430 261 181 100 | Assembly | 10 |
| 6 | Mounting insert | 430 261 185 000 | 4 pcs per scale | 10 |
| 7 | End cap Seal | 430 261 190 000 | 2 pcs per scale | 10 |
| 8 | Spring 4 | 430 261 186 000 | 2 pcs per scale | 10 |
| 9 | Dust lips | 430 261 245 100 |  | 20 m |
| 11 | Slider holder | 430 261 680 000 |  | 10 |

Note:	Make use of dust lips by cutting them with a cutter to the length needed for use.

| Accessory bracket parts, etc. |
| --- |
| No. | Part name | Part code | Remarks | Ordering unit |
| ⑦ | HSB M3x8 | 001 715 109 000 | Fixing Mounting el. to scale | 10 |
| ⑦ | PT counter-sunk screw KA25x6 | 001 717 115 000 | Fixing End cap to scale | 10 |
| ⑦ | HSB M8x20 | 001 715 120 000 | Fixing Scale to machine | 10 |
| ⑦ | HSB M4x20 | 001 715 105 000 | Fix. Scan. unit to machine | 10 |
| ⑧ | Clamp for wires small size | 001 981 230 000 | Fixing cable to machine | 10 |
| ⑧ | +P screw M4x10 | 001 710 356 000 | For clamp | 10 |
| ⑧ | Air inlet M5/(6x1 | 001 765 950 000 |  | 10 |
| ⑧ | User manual | 430 261 710 000 |  | 1 |

Note: Using parts in place of the parts determined by this company can lead to secondary malfunctions, and so should be avoided.

2-2. List of tools and supplies used

Prepare the following tools for making repairs to the TGM133 series.

# Special tools

| Tools for inserting Reading head |
| --- |
| Part name | Part code | Remarks | Ordering unit | Stock |
| Securing plate | 430 261 128 000 | For inserting Reading head | 10 | 0 |
| Adjustment jig | 430 261 814 000 |  | 1 | 0 |

# Adhesive agents

| Adhesive agent |
| --- |
| Part name | Part code | Remarks | Ordering unit | Stock |
| Silicone adhesive TERMOSIL K6 | 002 040 322 001 | For securing scale End cap | 1 | 0 |
| Loctite 222 | 002 040 154 000 | For securing screws | 1 | 0 |
| Silicone grease MERCK 1.07746. | 002 030 618 000 | For Dust lips, gaskets etc. | 1 | 0 |
| Service set of Silicone adhesive | 430 261 803 000 | For Dust lips, gaskets etc. | 1 | 0 |
| Service set of Loctite 222 | 430 261 802 000 | For Dust lips, gaskets etc. | 1 | 0 |
| Service set of Silicone grease | 430 261 801 000 | For Dust lips, gaskets etc. | 1 | 0 |

# General tools

| Tool name | Usage |
| --- | --- |
| Dial indicator or an equivalent | Since this is used in realignment during reinstallation, make use of an item capable of readings up to 2µm. |
| Magnet stand | Used to secure the dial indicator. |
| L-type wrench (M4, M5, M8) | Used when removing the scanning unit. |
| Torque driver and hexagon bits for M4, M5, M8 | Since this is used when reinstalling the scale, prepare an item capable of torque settings. |
| Cutting pliers or similar product | Used when removing dust lips, etc. |
| Cutters, nippers, etc. | Used to cut dust lips, etc. |
| Cotton swabs or a similar product | Used when cleaning the scale’s interior. |
| Ethyl alcohol or acetone | Used when cleaning the scale. |
| Oscilloscope | Since this is used to make waveform adjustments, prepare an item capable of observing resurge waveforms. |
| + Screwdriver | Make use of an item that can turn M2.5 screws. |
| Minus screwdriver | Since this is used to turn the adjustment volume, make use of small scale clock screwdrivers, etc. |
| Waste/duster or rags | Used to wipe up oil from around the scale. |
| Other items | Tweezers, (100 vinyl tubes, vinyl ties, etc. |